## Slide 1
Introduction
Photodynamic Therapy Delivery System
MEMBERS:
Michael Bernatzky, Cobi Ben-David, Ryan Donovan, Simon Ioffe, Andrew Zamsky
ADVISOR:
Professor B.K. Jaeger
CONSULTANT:
Professor A. Gouldstone
Notes: Say what the team is commissioned with right away

## Slide 2
40 - 50 million Americans
Acne
Up to 58 million Americans
Actinic Keratosis
US Population
312 million
13 million Americans
Skin Cancer
15 million Americans
Jaundice
Why Does This Matter?
Scope
Treatment
Definition

## Slide 3
40 - 50 million Americans
Acne
Up to 58 million Americans
Actinic Keratosis
US Population
312 million
13 million Americans
Skin Cancer
15 million Americans
Jaundice
Why Does This Matter?
What can be done?
Photosensitizing Agent (PSA)
Blue Light Delivery
Treatment of
Dermatological
Conditions
Photodynamic Therapy
Scope
Treatment
Definition

## Slide 4
What is PDT?
Photodynamic Therapy
PDT is a form of light-activated treatment, which works in conjunction with a photo-sensitizing agent (PSA) to target and destroy unhealthy or dangerous cells.
Photon
Activated Pharmaceutical
Pharmaceutical
O2
Oxidizing Species
(Singlet Oxygen)
Oxidative Injury to Cells
Cellular Death
Scope
Treatment
Definition
P. Acnes
PDT
P. Acnes Absent

## Slide 5
Current State
Problem Statement
Existing Devices
For Clinicians:
Current devices are difficult to lift, adjust and operate
The Problem
Existing devices are expensive which hinders their availability

## Slide 6
Current State
Problem Statement
Existing Devices
For Patients:
Current devices are hard to adjust to different parts of the body and various
body types
Treatments are often uncomfortable, dampening a patients desire to return
for treatment
The Problem

## Slide 7
Design Goals | 
Flexibility | ✗
User Interface | ✗
Ergonomic | ✗
Cost | !
Patient Comfort | ✗
Problem Statement
Existing Devices
DUSA®
Current State
✓
✗
Adequate
!
Poor
Excellent

## Slide 8
Design Goals |  | 
Flexibility | ✗ | !
User Interface | ✗ | ✗
Ergonomic | ✗ | !
Cost | ! | ✗
Patient Comfort | ✗ | !
Problem Statement
Existing Devices
Omnilux®
Current State
✓
✗
Adequate
!
Poor
Excellent

## Slide 9
The team has submitted an Institutional Review Board (IRB) application and received approval
All team members have successfully been certified by the National Institute of Health (NIH)
Preparation
Literature
Consultations
Patents
Research

## Slide 10
Research
Preparation
Literature
Consultations
Patents

## Slide 11
Research
Preparation
Literature
Consultations
Patents

## Slide 12
Research
Preparation
Literature
Consultations
Patents
M. Kaminer, MD
K. Nolan, RN, MSN
C. DiMarzio, PhD
Skincare Physicians
Dermatology Associates
Northeastern University Optics Lab

## Slide 13
Ease of use for clinicians
Flexible design for various treatment applications
Optimal patient comfort
Inexpensive cost to broaden accessibility
Responsibilities
Initial Concepts
Decision
Design Criteria
Design Process

## Slide 14
Industrial Engineering
Mechanical Engineering
Design Process
Responsibilities
Initial Concepts
Decision
Usability
Anthropometry
Ergonomics
Device Interface
Functionality
Material Selection
Cost Effectiveness
Patient Comfort
Flexible Optics
Modularity
Packaging
Adjustability
Design Criteria

## Slide 15
Design Process
Responsibilities
Initial Concepts
Decision
Design Criteria

## Slide 16
Design Process
Responsibilities
Initial Concepts
Design Criteria
Decision
Decision Matrix:

## Slide 17
Design Process
Responsibilities
Initial Concepts
Design Criteria
Decision
Halogen
LED
Fiber Optics
Total Score: 1900
Total Score: 2185
Total Score: 2625
Decision Matrix:

## Slide 18
Design Analysis
Electricity
Heat
Light
Fiber Optic Cables
Light Source
Patient
Flexible Light Mat
Light
Light
Concept Overview:
Overview
Treatment Area
MedLiTx
Wavelength Selection

## Slide 19
Design Analysis
Overview
Treatment Area
MedLiTx
Wavelength Selection
Notes: Say what the team is commissioned with right away

## Slide 20
Flexible Silicone Light Mat
Light Source
Flexible Fiber Optic Transmission Cable
Modular Cable Coupling
Design Analysis
Overview
Treatment Area
Wavelength Selection
MedLiTx

## Slide 21
Visible Light Spectrum
What wavelength light to use?
Most current PDT devices use one single light wavelength, while different treatments require different light at different wavelengths.
Our device incorporates multiple frequencies in a single device.
Blue Light Therapy
(primary design objective)
Used for other forms of PDT
(expanded design objective)
Safe
Effective
Approved
þ
þ
þ
Overview
Treatment Area
Wavelength Selection
Design Analysis
MedLiTx

## Slide 22
Light Source
Light Filter Array
Light Filter Array
The design gives clinicians flexibility by incorporating light at various wavelengths in a single device for different treatments.
Patient
White Light (380-750nm)
Blue light(450-475nm)
Violet light(380-450nm)
Red light(620-750nm)
Design Analysis
Overview
Treatment Area
Wavelength Selection
MedLiTx

## Slide 23
Face
Arms
Hands
Legs
Chest
Back
Torso
Fitting the Patient
Theses are the physical areas
of focus established from consultations with PDT practitioners and literature research.
A flexible design is important
to be able to fit all the parts
of the body comfortably.
Design Analysis
Overview
Treatment Area
Wavelength Selection
MedLiTx

## Slide 24
Head ≈ 9%
Chest ≈ 9%
Abdomen ≈ 9%
Upper/mid/low back and buttocks ≈ 18%
Each arm ≈ 9% (front ≈ 4.5%, back ≈ 4.5%)
Each leg ≈ 18% total (front ≈ 9%, back ≈ 9%)
The Rule of 9
No more than 9% of the body should be treated by PDT at a time.
Design Analysis
Overview
Treatment Area
Wavelength Selection
MedLiTx
Notes: Our devices treatment area 5.78%

## Slide 25
Light Delivery Mat
Rubber Silicon
(Flexible, Biocompatible)
Titanium Dioxide
(Light scattering agent)
Stainless Steel Mesh
(For ductility, malleability)
Reflective Metlar Sheet
(Reflects light forward)
Side-Glow Fiber Optic Cables
(Emits light throughout mat)
þ
Flexible
Biocompatible
Inexpensive
þ
þ
Sterilize with Ease
þ
Design Features
Modularity
Light Weight
Flexibility
User Interface
Notes: Say what the team is commissioned with right away

## Slide 26
Interchangeable Delivery System
We designed a fiber optic coupler that allows different delivery mechanisms to be plugged in, based on the specific treatment.
Physical Consideration
To mitigate losses when connecting fiber optic cables, the device is designed to use cables with the same numerical aperture.
Modularity
Light Weight
Flexibility
Design Features
User Interface

## Slide 27
Height ≈ 72 in
Load Constant ≈ 10Lb
Coupling Multiplier= 1
Vertical Location ≈ 36 in
Asymmetry Angle ≈ 140
Horizontal Multiplier=7.2
Frequency Multiplier=.95
Vertical Travel Distance ≈ 36 in
Distance Multiplier=.82+(1.8/36)=.87
Asymmetry Multiplier=1-.0032*140=.552
Vertical Multiplier= 1-(.003*[36-75])=1.117
NIOSH Lifting Limit Formula
Based on the design specifications the Recommended Weight Limit is:
Design Features
Modularity
Light Weight
Flexibility
User Interface

## Slide 28
Existing User Interface:
?
?
Huh?
What?
Design Features
Modularity
Light Weight
Flexibility
User Interface

## Slide 29
Design Features
Modularity
Light Weight
Flexibility
User Interface

## Slide 30
Authorization Screen
Treatment Setup
Paused Treatment
Treatment Status
Saved Programs
Program
Edit
Wavelength Edit Screen
Denied
Approved
Begin Treatment
Resume
Pause
Stop Treatment
Stop Treatment
Change
Use Program
Save Program
Edit Program
New Program
Name Edit Screen
Intensity Edit Screen
Time Edit Screen
Edit Parameter
Edit Parameter
Edit Parameter
Edit Parameter
Save Parameter
Save Parameter
Save Parameter
Design Features
Modularity
Light Weight
Flexibility
User Interface

## Slide 31
Design Features
Modularity
Light Weight
Flexibility
User Interface

## Slide 32
Device Comparison
Design Goals |  |  | 
Flexibility | ✗ | ! | ✓
User Interface | ✗ | ✗ | ✓
Ergonomic | ✗ | ! | ✓
Cost | ! | ✗ | ✓
Patient Comfort | ✗ | ! | ✓
✓
✗
Adequate
!
Poor
Excellent
Breast Cancer Tx | ! | ✗ | ✓
Design Features

## Slide 33
Thank you for listening.
Q&A
Questions ?

## Slide 34
Appendices
IE Contributions
User Interface Flowchart
Design Criteria
Photodynamic Therapy Chemical Process
Responsibilities
Light Mechanism / Treatment Safety
ME Contributions
Prototype Manufacturing Process
List of Terms(A-H)
List of Terms(I-Z)
Prototype Costs

## Slide 35
Appendix
IE Contributions
Anthropometry & Ergonomics:
Usability:
Manufacturing:
Appendix

## Slide 36
Appendix
ME Contributions
Material Selection:
Rubber Silicon
Stainless Steel Mesh
Reflective Metlar Sheet
Side-Glow Fiber Optic Cables
Device Design:
Appendix
Flexible Silicone Light Mat
Light Source
Flexible Fiber Optic Transmission Cable
Modular Cable Coupling
Titanium Dioxide

## Slide 37
Appendix
Light Mechanism/ Treatment Safety
Appendix
Safe
Effective
Approved
þ
þ
þ
Photosensitizing Agent (PSA)
Light Delivery
Treatment of
Dermatological
Conditions
Visible Light Spectrum
Blue Light Therapy
(primary design objective)
Used for other forms of PDT
(expanded design objective)

## Slide 38
Appendix
User Interface Flowchart
Appendix
Authorization Screen
Treatment Setup
Paused Treatment
Treatment Status
Saved Programs
Program
Edit
Wavelength Edit Screen
Denied
Approved
Begin Treatment
Resume
Pause
Stop Treatment
Stop Treatment
Change
Use Program
Save Program
Edit Program
New Program
Name Edit Screen
Intensity Edit Screen
Time Edit Screen
Edit Parameter
Edit Parameter
Edit Parameter
Edit Parameter
Save Parameter
Save Parameter
Save Parameter

## Slide 39
Optimal patient comfort
Ease of use for clinicians
Inexpensive cost to broaden accessibility
Flexible design for various treatment applications
Design Criteria
Appendix
Appendix

## Slide 40
Photodynamic Therapy
PDT is a form of light-activated treatment, which works in conjunction with a photo-sensitizing agent (PSA) to target and destroy unhealthy or dangerous cells.
Photon
Activated Pharmaceutical
Pharmaceutical
O2
Oxidizing Species
(Singlet Oxygen)
Oxidative Injury to Cells
Cellular Death
Appendix
Photodynamic Therapy Chemical Process
Appendix

## Slide 41
Industrial Engineering
Mechanical Engineering
Usability
Anthropometry
Ergonomics
Device Interface
Functionality
Material Selection
Cost Effectiveness
Patient Comfort
Flexible Optics
Modularity
Packaging
Adjustability
Appendix
Appendix
Responsibilities

## Slide 42
Appendix
Appendix
Manufacturing Process

## Slide 43
Appendix
Appendix
Prototype Costs
Parts | Unit Cost | QTY | Total Cost
2 mm Solid Core: Side Glow Fiber | $2.75/ft | 40 | $110
125 grams Titanium Dioxide | $4.30 | 1 | $4.30
2 mil Reflective Metlar | $2 | 1 | $2.00
Moldmax 27T Silicone 1 Gallon | $88.82 | 1 | $88.82
5mm Solid Core End glow Fiber | $3./ft | 70 | $210
150 Watt Metal Halide Illuminator w/8 color wheel | $400 | 1 | $400
Illuminator-End glow fiber Coupler | $60 | 1 | $0
End glow-Side glow fiber Coupler | $60 | 1 | $0
Silicone Mold | $180 | 1 | $0
 | Prototype Parts Cost: |  | $815.12
Additional Manufacturing Costs |  |  | 
Coupler Machining | $60/hr | 2 | $120
Mold Machining | $60/hr | 3 | $180
Silicone Molding | $60/hr | 3 | $60
Assembly | $60/hr | 3 | $60
 | Additional Mfg. Cost: |  | $420
 |  | Total Cost | $1,235

## Slide 44
Appendix
Appendix
List of Terms (A-H)
Actinic Keratosis: An overgrowth of skin layers resulting from extended exposure to the sun
Analgesia: A deadening or absence of the sense of pain without loss of consciousness
Apoptosis: a type of cell death in which the cell uses specialized cellular machinery to kill itself
Biocompatible: Not harmful to living tissue
Cladding: A covering or coating on a structure or material
Claustrophobia: Extreme or irrational fear of confined places
Coagulation necrosis: A type of accidental cell death typically caused by ischemia or infarction. This is in contrast to apoptosis, which is a naturally occurring cause of cellular death. While apoptosis often provides beneficial effects to the organism, necrosis is almost always detrimental and can be fatal.
Collimator: A device for producing a parallel beam of rays or radiation
Dermatology: The branch of medicine concerned with the diagnosis and treatment of skin disorders
Dosiology: The study and determination of the doses in which medicines should be administered
Dosimetry: The accurate measurement of doses, especially of radiation
Drug-to-light interval: The period of time between when the drug is given and when the light is applied
Epithelial tissue: Epithelium; membranous tissue covering internal organs and other internal surfaces of the body
Facilitator: A device, tool, or person that helps to bring an outcome by providing indirect or unobtrusive assistance, guidance, or supervision
Hysteresis: The phenomenon in which the value of a physical property lags behind changes in the effect causing it

## Slide 45
Appendix
Appendix
Illuminator: A device for producing, concentrating, or reflecting light
Intravenous: Existing within, or administered into, a vein or veins.
Irradiation: Exposure to radiant energy (heat, x-rays, etc.) for therapeutic or diagnostic purposes
Jaundice: A medical condition with yellowing of the skin or whites of the eyes, arising from excess of the pigment bilirubin and typically caused by an accumulation of bile pigment
Levulan: Topical medication manufactured by DUSA, to be used in conjunction with their blue light to treat acne
Metvixia: Topical medication often administered along with PDT to make skin more sensitive to light
Noninvasive: Not penetrating the body, as by incision or injection
Oncology: The study of Cancer. The branch of medical science dealing with tumors, including the origin, development, diagnosis, and treatment of malignant neoplasms
Palliative Nursing: form of medical care or treatment that concentrates on reducing the severity of diease symptoms, rather than striving to halt, delay or reverse progression of the disease itself or provide a cure
Pathology: The Science of the causes and effects of diseases, the typical behavior of a disease
Photosensitization: skin condition marked by heightened sensitivity to artificial and natural light
Photosensitizer: A light-absorbing substance that initiates a photochemical or photo-physical reaction in another substance (molecule), and is not consumed in the reaction
Photodynamic Therapy (PDT): A form of light-activated therapy that may use oral or topical photosensitizing agents (PSA)
Spectrometer: An instrument used for measuring wavelengths of light
Sub-dermally: beneath the skin
Topically: Locally; to a restricted area of the body
Zimmer: the air hose used in light therapy to minimize burning sensation caused by light using cold air
List of Terms (I-Z)